## INSTRUCTIONS
SWATH Performance Kit Performance Tracker
The SWATH Performance Kit and Protocol has been specifically designed to enable users to evaluate and optimize the performance of their LC-MS systems. Instructions and methods for performing the tests on TripleTOF® systems can be found in the SWATH Performance Kit Protocol. This also provides a great way for new users to get up and running quickly doing these proteomics workflows.
To assess the performance of the LC-MS systems, four system suitability tests are available.
First, the PepCalMix Infusion Test is performed to check the MS performance and to optimize the source
positioning and parameters.
Next, the PepCalMix LC-MS System Suitability Test is performed to make sure that the system is in basic working order and is providing good LC separation and MS sensitivity. This method is also used for instrument mass calibration.
Next, the Cell Lysate SWATH® Acquisition QC Test is performed to assess the performance of the SWATH Acquisition and includes a digested human cell lysate with PepCalMix added.
Finally, the Cell Lysate IDA QC Test is performed to assess the protein identification performance.
This Excel Performance tracker contains the performance guidelines that users can use to assess the performance status of their individual systems and to monitor their system over time to ensure consistent data quality.
Performance on an individual MS system can be affected by many different variables including (but not limited to) instrument to instrument variability, instrument tuning, source optimization differences, trap and column variability and age, etc. To use these tests for performance benchmarking, it is the user responsibility to follow the instructions in the Protocol as closely as possible using the configuration defined. Troubleshooting tips have been provided in the Proteomics User community to help users obtain their results. However, because of the complexity and variability possible with these tests, SCIEX cannot guarantee performance of these tests.
For Research Use Only. Not for use in diagnostic procedures. Not for use in clinical investigational procedures.

## TripleTOF 6600 micro
TripleTOF 6600 Micro - Performance Assessment
| See SWATH Acquisition Performance Kit Protocol for instructions |  |  | Date
Name
Instrument (serial #)
PepCal Lot Number
Cell Lysate Lot Number
Instrument Cleaned Before (Y/N)
Instrument Tuned Before (Y/N)
New Parts Installed on LC (column, trap, AS needle..)
Recommendation = Performance Threshold
Test 1: PepCalMix Infusion Test - Source Tuning
50 fmol/uL PepCalMix - 5 uL/min flow rate
TOF MS High Resolution (1 sec accum) |  | Peak m/z
| Sum Intensity | 533.3233 |  | 15000
| Sum Intensity | 613.3495 |  | 13000
| Sum Intensity | 677.853 |  | 7000
| Sum Intensity | 758.9105 |  | 7000
| Sum Intensity | 964.9771 |  | 30000
| TOF MS Resolution | 533.3233 |  | 25000
| TOF MS Resolution | 613.3495 |  | 26000
| TOF MS Resolution | 677.853 |  | 26000
| TOF MS Resolution | 768.9034 |  | 27000
| TOF MS Resolution | 964.9771 |  | 30000
TOF MS/MS 964.98 High Resolution (1 sec accum) (CE 45_5CES)
| Sum Intensity | 1252.6043 |  | 4000
| Sum Intensity | 596.339 |  | 700
| Peak Resolution | 1252.6043 |  | 26000
| Peak Resolution | 596.339 |  | 23000
TOF MS/MS 964.98 High Sensitivity (1 sec accum) (CE 45_5CES)
| Sum Intensity | 1252.6043 |  | 10000
| Sum Intensity | 596.339 |  | 1200
| Peak Resolution | 1252.6043 |  | 15000
| Peak Resolution | 596.339 |  | 14000
TOF MS/MS 758.9 High Resolution (1 sec accum) (CE 40_5CES)
| Sum Intensity | 1070.5972 |  | 250
| Sum Intensity | 348.1554 |  | 250
| Peak Resolution | 1070.5972 |  | 26000
| Peak Resolution | 348.1554 |  | 23000
TOF MS/MS 758.9 High Sensitivity (1 sec accum) (CE 40_5CES)
| Sum Intensity | 1070.5972 |  | 500
| Sum Intensity | 348.1554 |  | 500
| Peak Resolution | 1070.5972 |  | 15000
| Peak Resolution | 348.1554 |  | 14000
TOF MS/MS 533.3 High Resolution (1 sec accum) (CE 25_5CES)
| Sum Intensity | 824.4967 |  | 700
| Sum Intensity | 355.234 |  | 400
| Peak Resolution | 824.4967 |  | 26000
| Peak Resolution | 355.234 |  | 23000
TOF MS/MS 533.3 High Sensitivity (1 sec accum) (CE 25_5CES)
| Sum Intensity | 824.4967 |  | 1400
| Sum Intensity | 355.234 |  | 800
| Peak Resolution | 824.4967 |  | 15000
| Peak Resolution | 355.234 |  | 14000
Troubleshooting – Lower than Desired MS Sensitivity or Resolution
Troubleshooting – Lower than Desired MS/MS Sensitivity or Resolution
Test 2: PepCalMix LC-MS System Suitability Test
40 fmol PepCalMix on column
TOF MS - 250 msec - 0.05Da XIC peak area |  | Peak m/z | Approx Peak RT | Peak Area
| AETSELHTSLK | 408.5501 | 5.26 | 250000
| GAYVEVTAK | 473.2602 | 5.5 | 250000
| IGNEQGVSR | 485.253 | 4.58 | 200000
| LVGTPAEER | 491.2656 | 5.16 | 300000
| LDSTSIPVAK | 519.7997 | 5.77 | 300000
| AGLIVAEGVTK | 533.3233 | 6.38 | 400000
| LGLDFDSFR | 540.2734 | 7.83 | 200000
| GFTAYYIPR | 549.2863 | 7.03 | 200000
| SGGLLWQLVR | 569.834 | 8.1 | 70000
| AVGANPEQLTR | 583.3136 | 5.61 | 400000
| SAEGLDASASLR | 593.8005 | 5.89 | 250000
| VFTPLEVDVAK | 613.3495 | 7.25 | 400000
| VGNEIQYVALR | 636.3527 | 6.71 | 400000
| YIELAPGVDNSK | 657.345 | 6.41 | 400000
| DGTFAVDGPGVIAK | 677.8583 | 6.67 | 200000
| YDSINNTEVSGIR | 739.3615 | 6.01 | 150000
| SPYVITGPGVVEYK | 758.9105 | 6.78 | 200000
| ALENDIGVPSDATVK | 768.9034 | 6.33 | 500000
| AVYFYAPQIPLYANK | 883.4738 | 7.81 | 200000
| TVESLFPEEAETPGSAVR | 964.9771 | 7.18 | 1000000
|  |  | Median of 20 Peaks | 250000
|  |  | # of Peptides Passing | 16
High Sensitivity MS/MS - 500 msec - 0.05Da XIC peak area
| TOF MS/MS XIC peak area | 758.91 / 1070.6 |  | 30000
High Resolution MS/MS - 500 msec - 0.05Da XIC peak area
| TOF MS/MS XIC peak area | 758.91 / 1070.6 |  | 15000
Chromatography Quality
| Peak width half height (min) |  |  | 0.07
| RT spread 485 m/z to 569 m/z (min) |  |  | 3
Troubleshooting - Loss of Early Eluting Peptides During Trap-Elute Workflow
Troubleshooting – Spread of Peptide Elution is Compressed
Troubleshooting –Peptide Elution is Delayed
Troubleshooting – LC-MS Peak Area is Lower than Expected
Test 3: Cell Lysate SWATH® Acquisition QC Test
Cell Lysate (4ug) + PepCal Mix (400fmol)
SWATH® Acquisition 2.0 MicroApp Results - <1% FDR filter |  |  | #
| SWATH Replicate Template - Proteins with %CV < 20% |  | Proteins | 2400
| SWATH Replicate Template - Peptides with %CV < 20% |  | Peptides | 9500
Troubleshooting - Loss of Early Eluting Peptides During Trap-Elute Workflow
Test 4: Cell Lysate IDA QC Test
Cell Lysate IDA (4ug)
|  |  | # of MS/MS spectra | 50000
ProteinPilot™ Software Results
| Protein | Local | 1%
| Protein |  | 5%
| Protein |  | 10%
| Protein | Global | 1% | 2300
| Protein |  | 5%
| Protein |  | 10%
| Distinct peptide | Local | 1%
| Distinct peptide |  | 5%
| Distinct peptide |  | 10%
| Distinct peptide | Global | 1% | 15000
| Distinct peptide |  | 5%
| Distinct peptide |  | 10%
| Spectral | Local | 1%
| Spectral |  | 5%
| Spectral |  | 10%
| Spectral | Global | 1% | 25000
| Spectral |  | 5%
| Spectral |  | 10%
Troubleshooting - Loss of Early Eluting Peptides During Trap-Elute Workflow